## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E2330-19
ANSI/ASTM E2330-19
Standard Test Method for Determination of Concentrations of Elements in Glass Samples Using Inductively Coupled Plasma Mass Spectrometry (ICP-MS) for Forensic Comparisons
Discipline: | Trace Evidence
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Apparatus | 5 | Section Title
ICP-MS | 5.1 | Recommendation | An Inductively Coupled Plasma Mass Spectrometer (ICP-MS) instrument is employed. Since there are several manufacturers for ICP-MS instruments, the instrument maker, model, conﬁguration, and major operational parameters should be noted within the analysis results.
Standard Reference Glasses | 5.2 | Recommendation | A minimum of two different standard reference glasses of known elemental composition should be used. Examples suitable for this analysis include NIST 1831 and NIST 612 Reference Glasses
Non-Glass Laboratory Ware | 5.3 | Section Title
Micro-Balance | 5.4 | Informational Text | A microbalance with a precision of 61 µg or better is necessary for the accurate weighing of the glass samples.
High Purity Reagents | 5.5 | Requirement | ICP-MS grade acids and reagents shall be used for digestion and dilution.
Apparatus | 5.6 | Informational Text | Laboratory Oven or Dry Bath Block.
Apparatus | 5.7 | Informational Text | Micropipettes.
Apparatus | 5.8 | Informational Text | Fume Hood.
Hazards | 6 | Section Title
Hazards | 6.1 | Requirement | Hydroﬂuoric acid (HF) is corrosive to all body tissues and bone. Skin, eye or respiratory contact with HF presents a signiﬁcant health hazard that requires immediate ﬁrst aid and possibly specialized medical care. Extreme care must be exercised in the presence of this acid. Calcium gluconate gel should be kept in the immediate area of HF use in the event of a potential skin exposure. See the chemical manufacturer’s Safety Data Sheet for more information.
Hazards | 6.2 | Recommendation | The argon plasma should not be observed directly without protective eyewear. Potentially hazardous UV light may be emitted.
Hazards | 6.3 | Recommendation | ICP-MS instruments generate high amounts of radio-frequency energy in their RF power supply and torch boxes that is potentially hazardous if allowed to escape. Safety devices and safety interlocks should not be bypassed or disconnected.
Sample Preparation | 7 | Section Title
Sample Preparation | 7.1 | Requirement | The sample set for analysis shall include all known samples, questioned samples and at least two standard reference glasses. Prior to crushing the glass sample for the digestion, soak samples in concentrated HNO3, rinse 3 times with high purity water, and allow the samples to dry.
Sample Preparation | 7.2 | Requirement | Crush the samples between clean polymeric materials, such as polystyrene weighing boats or glassine sheets, taking care not to puncture the materials.
Sample Preparation | 7.3 | Requirement | Accurately weigh approximately 2 to 3 mg of each sample using a microbalance (with a precision of ±1 μg or better) and quantitatively transfer into a labeled non-glass tube with a cap. At least three weighings per glass source shall be made for a minimum of three analytical samples per glass source for digestion. Prepare empty labeled non-glass tubes for reagent blanks.
Sample Preparation | 7.4 | Requirement | All volumes are delivered using micropipets. Add concentrated hydrofluoric acid, concentrated hydrochloric acid, and concentrated nitric acid to each tube to make a 2 to 1 to 1 mixture of the acids in the tubes.
Sample Preparation | 7.5 | Requirement | Cap and vortex mix the tubes. The tubes may be placed in an ultrasonic bath to assist in the digestion for approximately one hour. Uncap the tubes and place in a dry bath block or an oven, at 80°C or greater (but below the softening temperature of the digestion tubes), until dry.
Sample Preparation | 7.6 | Requirement | Reconstitute the samples by adding 500 μL of 50 % HNO3 (8.0 mol/L). Recap the tubes and vortex mix.
Sample Preparation | 7.7 | Informational Text | The tubes may be placed in an ultrasonic bath for at least one hour or left to stand overnight.
Sample Preparation | 7.8 | Requirement | Add 50 μL of a 10 mg/L Rh internal standard solution and 4450 μL of ultrapure water to each tube and vortex mix contents. Each tube will contain a 5 mL solution with 100 μg/L Rh internal standard in 5 % HNO3.
Instrument Set-Up And Calibration | 8 | Section Title
Instrument Set-Up And Calibration | 8.1 | Requirement | Tune the instrument prior to the analysis using the manufacturer’s recommendations covering the mass range of the identified elements. The instrument shall be adjusted for maximum sensitivity, best precision, and to minimize oxides and doubly charged ion interferences.
Instrument Set-Up And Calibration | 8.2 | Requirement | Prepare calibration standards from pure element standards traceable to accepted metrological sources (NIST, etc.) covering the expected range of concentrations of the glass samples.
Instrument Set-Up And Calibration | 8.3 | Requirement | Prepare two calibration curves as well as two check standards. The first calibration curve shall consist of 24Mg, 27Al, 47Ti, 57Fe, 55Mn, 88Sr, 90Zr, 138Ba, and 206, 207, 208Pb with a concentration range of 0.0, 1.0, 10.0, 50.0, 75.0, and 150.0 μg/L. The second calibration curve shall consist of 85Rb, 139La, 140Ce, 146Nd, 148Sm with a concentration range of 0.0, 0.1, 0.5, 1.0, 5.0, and 50 μg/L. An internal standard of 100 μg/L Rh shall be used in each standard sample.
Instrument Set-Up And Calibration | 8.4 | Requirement | The check standard (continuing calibration verification or CCV) for the element standards calibration shall be 50.0 μg/L for the first group and 5.0 μg/L for the second group.
Instrument Set-Up And Calibration | 8.5 | Requirement | Analyze the standard samples using the ICP-MS and calibration curves established for each group of elements. Analyze the continuing calibration verification (CCV) samples. The system should be recalibrated any time that the CCV falls outside the acceptable parameters established by the laboratory or analyst for this procedure.
Sample Analysis | 9 | Section Title
Sample Analysis | 9.1 | Requirement | Analyze a reagent blank with every sample set.
Sample Analysis | 9.2 | Requirement | Analyze blanks between replicate groups.
Sample Analysis | 9.3 | Requirement | Analyze two Standard Reference Glasses (NIST 1831, NIST 612, etc.) as a part of the sample set.
Sample Analysis | 9.4 | Requirement | Analyze all blanks and digested glass samples using the instrument parameters established during tuning. The data are compared to the calibration curves, concentration values are determined, and corrected for the weight of the glass sample used. Final concentration values should be reported as milligram of element per kilogram of glass (mg/kg).
Quality Control | 10 | Section Title
Continuing Calibration Verification (CCV) | 10.1 | Recommendation | After measuring the standards and periodically throughout the analysis, a CCV sample is measured to determine if the calibration is valid. The CCV sample is another solution containing all of the elements of interest at a known concentration (50.0 μg/L and 5.0 μg/L). The system is recalibrated any time one or more element concentrations measured for the CCV sample fall outside the acceptable concentration ranges for the actual concentration of the elements in the CCV. The acceptable ranges of the concentrations of the elements in the CCV and the frequency of CCV sample measurements should be determined by each laboratory using this method.
Method Detection Limit (MDL) and Limit of Quantitation (LOQ) | 10.2 | Informational Text | The MDL for each element is determined by measuring the concentrations of the elements of interest in three procedure blanks on two non-consecutive days. The MDL is the standard deviation of the concentration of each element in the three procedure blanks times three. The LOQ is the standard deviation of the concentration of each element in the three procedure blanks times ten. To calculate these limits of detection and quantitation, the average from the results for the two days is taken. NOTE 1: Table X2.1 illustrates the values for MDL, LOQ, and range of concentrations found in a set of 50 soda-lime glass samples all expressed as ug/L of the elemental concentration in the final solution.
Accuracy of Standard Reference Material (SRM) | 10.3 | Recommendation | SRMs with known concentrations of elements in glass are measured to assess the accuracy of the method. If one or more element concentrations fall outside of the acceptable ranges of concentrations (as determined by each laboratory), the sample analysis can yield inaccurate concentrations. Corrective actions should be taken to address the source of the inaccuracy prior to additional analysis.
Calculation And Interpretation of Results | 11 | Section Title
Calculation And Interpretation of Results | 11.1 | Requirement | The procedure below shall be followed to conduct a forensic glass comparison when using the recommended match criteria.
Calculation And Interpretation of Results | 11.1.1 | Requirement | For the Known source fragments, using a minimum of 3 measurements, calculate the mean for each element.
Calculation And Interpretation of Results | 11.1.2 | Requirement | Calculate the standard deviation for each element. This is the Measured SD.
Calculation And Interpretation of Results | 11.1.3 | Requirement | Calculate a value equal to 3 % of the mean for each element. This is the Minimum SD.
Calculation And Interpretation of Results | 11.1.4 | Requirement | Calculate a match interval for each element with a lower limit equal to the mean minus 4 times the SD (Measured or Minimum, whichever is greater) and an upper limit equal to the mean plus 4 times the SD (Measured or Minimum, whichever is greater).
Calculation And Interpretation of Results | 11.1.5 | Requirement | For each Recovered fragment, using a minimum of 3 measurements, calculate the mean concentration for each element.
Calculation And Interpretation of Results | 11.1.6 | Requirement | For each element, compare the mean concentration in the Recovered fragment to the match interval for the corresponding element from the Known fragments.
Calculation And Interpretation of Results | 11.1.7 | Informational Text | If the mean concentration of one (or more) element(s) in the Recovered fragment falls outside the match interval for the corresponding element in the Known fragments, the element(s) does not "match" and the glass samples are considered distinguishable.
Calculation And Interpretation of Results | 11.2 | Informational Text | The expected error rates when using the recommended criteria for the comparison of elemental composition of glass are reported in the literature. However, the method is not meant to be exclusive, and other methods that provide similar or better performance in terms of error rates than the recommended method can be used in the practice.
Precision And Bias | 12 | Section Title
Precision And Bias | 12.1 | Informational Text | An interlaboratory study was conducted in 2001. Each of four laboratories tested four standard reference glasses using 5 replicate sample measurements of NIST 612, NIST 614, NIST 621, and NIST 1831.
Precision And Bias | 12.2 | Informational Text | The bias and precision results for each of the glasses are tabulated in Appendix X3. The terms repeatability and reproducibility are used as specified in Practice E177. The 95 % limits were calculated by multiplying the respective standard deviations by a factor of 2.8, as described in Practice E177.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics