## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E2937-18
ANSI/ASTM E2937-18
Standard Guide for Using Infrared Spectroscopy in Forensic Paint Examinations
Discipline: | Trace Evidence
Checklist version: | V1.1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Sample Handling | 6 | Section Title
Sample Handling | 6.1 | Requirement | The general collection, handling, and tracking of samples shall meet or exceed the requirements of Practice E1492 as well as the relevant portions of the SWGMAT's Trace Evidence Quality Assurance Guidelines.
Sample Handling | 6.2 | Requirement | The work area and tools used for the preparation of samples shall be free of all extraneous materials that could transfer to the sample.
Sample Handling | 6.3 | Recommendation | As stated in Guide E1610, a paint specimen should first be examined using a stereomicroscope, noting its size, appearance, layer sequence, heterogeneity within any given layer, and presence of any material that could interfere with the analysis (for example, traces of adhesive, surface abrasion transfers, or zinc phosphate conversion coating residue on the underside of the base primer on electro-coated parts). Some surface materials can be of interest and therefore can be worthy of analysis before removal.
Sample Handling | 6.4 | Recommendation | Each layer of a multi-layered paint should be analyzed individually.
Sample Handling | 6.5 | Requirement | When analyzing difficult items (for example, smears, dirty or heterogeneous specimens) care shall be taken when sampling the paint and in choosing appropriate analytical conditions. An attempt should be made to remove any extraneous material from the exhibit before sampling. If possible, analyze a number of areas to ensure reproducibility and understand inter/intrasample compositional variation.
Sample Handling | 6.6 | Recommendation | Extraneous material should be removed either by scraping with a suitable tool such as a scalpel or washing with water. If needed, alcohols or light aliphatic hydrocarbons can be useful for cleaning. However, it should be noted that the use of organic solvents for cleaning paint can alter the composition by extracting soluble components such as plasticizers or dissolve the paint binder. Therefore, cleaning with organic solvents should be considered the method of last resort. If solvents are used, known and unknown samples should be treated the same, making sure no residual solvent remains.
Sample Handling | 6.7 | Recommendation | For comparison of paint evidence, samples should be prepared and analyzed in the same manner.
Analytical Techniques and Operating Conditions | 7 | Section Title
Analytical Techniques and Operating Conditions | 7.1 | Requirement | Paints can be analyzed by transmission or reflection methods using the microscope accessory or bench accessories. The technique chosen is dependent upon the physical nature of the paint, the quantity of sample, preparation and analysis time, available equipment, and access to reference libraries for that technique. The same technique should be used on both known and questioned samples. Use multiple preparation or analytical techniques, or both, as necessary to analyze all layers and characteristics.
Analytical Techniques and Operating Conditions | 7.2 | Informational Text | The type of detector and beam splitter dictates the spectral range of the FTIR spectrometer. Mid-range infrared instruments use alkali halide beam splitters that are made from either cesium iodide (CsI) or potassium bromide (KBr).
Analytical Techniques and Operating Conditions | 7.3 | Informational Text | The most common infrared detector used on the main bench is a deuterated triglycine sulfate (DTGS) detector. The DTGS detector operates at room temperature. A spectrometer equipped with a DTGS detector and CsI optics has an approximate spectral range of 4000 to 200 cm-1. With KBr optics and a DTGS detector, the spectral range of the spectrometer is approximately 4000 to 400 cm-1.
Analytical Techniques and Operating Conditions | 7.4 | Informational Text | The detector commonly used with the microscope accessory is a mercury-cadmium-telluride (MCT) detector. The MCT detector is approximately 40× more sensitive than the DTGS detector, but has a narrower spectral range with a lower limit of 700 to 450 cm-1, depending on the type.
Analytical Techniques and Operating Conditions | 7.5 | Informational Text | Infrared data are collected from both the sample and a previously stored or newly acquired background. Taking the ratio of the sample spectrum to the background enables removal of absorptions from the cell or support material (for example, diamond absorptions) or from the atmosphere (for example, carbon dioxide and water vapor), or both. The latter absorptions can be minimized by purging with dried and filtered air desiccant packs or nitrogen gas. The number of scans acquired for each specimen can vary depending on sample type and size.
Main Bench Transmission Techniques | 7.6 | Section Title
Main Bench Transmission Techniques | 7.6.1 | Informational Text | A common bench transmission technique for the analysis of paint is the use of the diamond cell with a beam condenser.
Main Bench Transmission Techniques | 7.6.1.1 | Informational Text | Either prior to or after analysis, a background spectrum of the empty diamond cell is collected. The same background spectrum can be used for multiple samples or a new one can be collected for each sample analysis.
Main Bench Transmission Techniques | 7.6.1.2 | Recommendation | A sample from a single paint layer is placed on the clean diamond anvil cell and compressed between the windows to a desired thickness. Both high-pressure and low-pressure diamond cells can be used in conjunction with a beam condenser. Sample compression is normally done under a stereomicroscope to ensure uniform coverage. The cell is then placed in the sample holder in the main bench of the instrument. The instrument is allowed to equilibrate. This process is dependent upon the type of instrument and efficiency of the purge. A spectrum is then collected with the sample in place. Typically 16 to 256 scans are collected with a resolution of 4 cm-1. These parameters can vary depending on the instrument and size and nature of the sample. The same instrumental parameters, including the number of scans, should be acquired for the background as for the sample.
Main Bench Transmission Techniques | 7.6.1.3 | Informational Text | Diamond absorbs infrared radiation in the 2300 to 1900 cm-1 region; therefore, sample absorptions in this region can be obscured if the diamond path length is too long.
Attenuated Total Reflectance (Main Bench) | 7.7 | Section Title
Attenuated Total Reflectance (Main Bench) | 7.7.1 | Informational Text | Numerous in-bench single reflection ATR accessories are commercially available. The general principles of operation are the same for each accessory. The sample of interest is placed in direct contact with the internal reflecting crystal, such as diamond or KRS-5. Some accessories employ a viewing microscope to facilitate proper placement of the sample or area of interest.
Attenuated Total Reflectance (Main Bench) | 7.7.2 | Informational Text | In contrast to transmission methods, ATR methods require little or no sample preparation, although the force applied when using the ATR accessory can deform the sample.
Attenuated Total Reflectance (Main Bench) | 7.7.3 | Informational Text | Once in contact with the crystal, multiple scans are collected. The material is then removed and the crystal is cleaned. Background scans are collected with the item removed, either before or after the sample scans. Typically 16 to 256 scans are collected at a resolution of 4 cm-1. These parameters can vary depending on the instrument and size and nature of the sample.
FTIR Microscope Accessory | 7.8 | Section Title
FTIR Microscope Accessory | 7.8.1 | Requirement | In forensic science, infrared microspectroscopy is commonly used for acquiring the infrared spectrum of a paint. Spectra can be obtained from samples as small as 10 to 20 μm in diameter, using transmission and reflection methods. MCT detectors are commonly used with microscopes due to the higher sensitivity needed for small samples. These detectors are available in configurations usually designated as narrow band and broad (wide) band with the lower energy cut-off ranging from approximately 700 to 450 cm-1. There is a compromise between sensitivity and spectral range with these detectors. A detector with the spectral range of 4000 to 650 cm-1 and an area of 0.25 × 0.25 mm2 is typically used for paint examination since it offers the optimal balance between spectral range and sensitivity. These detectors shall be cooled by liquid nitrogen before use. When using the lower sensitivity/broader spectral range detector, larger samples are required.
Transmission Measurements | 7.8.2 | Section Title
Transmission Measurements | 7.8.2.1 | Requirement | Transmission methods generally require more extensive sample preparation. The sample shall be thin enough not to over-absorb. For transmission data viewed in % transmittance, spectral peaks optimally should not fall below 10 % T. For spectra displayed in absorbance, the maximum absorbance optimally should be 1.0 or less.
Transmission Measurements | 7.8.2.2 | Informational Text | A prepared and mounted sample is placed on the microscope stage and focused. The condenser on some instruments can be adjusted to account for the thickness of a support window. The sample is observed with visible light and the area to be analyzed is centered in the field of view. The area of interest is isolated from the remainder of the field of view with one or two apertures.
Transmission Measurements | 7.8.2.3 | Informational Text | The number of apertures corresponds to the instrument configuration. The apertures control the area and location of the infrared beam striking the sample and the transmitted light reaching the detector.
Transmission Measurements | 7.8.2.4 | Informational Text | Apertures also block unwanted radiation originating outside of the area of interest. If stray light is allowed to reach the detector, absorption intensity is reduced.
Transmission Measurements | 7.8.2.5 | Informational Text | As a sample area of interest becomes smaller, or as the aperture(s) are reduced so that a portion of the sample can be isolated, diffraction effects rapidly increase. These effects can be experienced when using aperture sizes smaller than 25 μm × 25 μm.
Transmission Measurements | 7.8.2.6 | Recommendation | To minimize the effects of heterogeneity, aperture areas greater than 2500 μm2 (for example, 50 μm × 50 μm or 25 μm × 100 μm) should be used when possible. Alternatively, multiple areas of the sample can be analyzed to determine the range of spectral characteristics.
Transmission Measurements | 7.8.2.7 | Informational Text | Once the area of interest is isolated by adjusting the magnification and apertures of the microscope, the infrared spectrum of the sample is collected. Typically 16 to 256 scans are collected at a resolution of 4 cm-1.
Transmission Measurements | 7.8.2.8 | Informational Text | The background spectrum is collected from an unused area of the support window using the same aperture configuration as used for the sample.
Transmission Measurements | 7.8.2.9 | Informational Text | If sample size is limited, the resulting spectrum can be noisy. To increase the signal to noise ratio (S/N), the number of scans can be increased. It is important to collect spectra with high S/N to permit visualization of fine detail such as small sharp peaks or shoulders in the resultant spectrum.
Reflection Measurements (Microscope) | 7.8.3 | Section Title
Reflection Measurements (Microscope) | 7.8.3.1 | Informational Text | The FTIR can also be used in the reflectance mode. However, most of the reference IR spectral data of coatings, binders, pigments and additives consist of transmission spectra. Furthermore, being surface analysis techniques, inconsistencies in the preparation of surfaces subject to comparative assessments can be problematic for data interpretation. Additionally, when analyzing individual layers in cross section and using the requisite small apertures, signal-to-noise constraints are even greater than those encountered in transmission analyses.
Reflection/Transmission (Transflection) | 7.8.4 | Section Title
Reflection/Transmission (Transflection) | 7.8.4.1 | Informational Text | This method is used to analyze thin films of a sample on an infrared reflective surface, such as a glass slide made of infrared light reflecting architectural glass (low e-glass), or metal. The sample is viewed using visible light and the area to be analyzed is centered in the field of view. The area of interest is isolated from the remainder of the field of view with an aperture and the infrared spectrum is collected. Typically, 16 to 256 scans are collected at a resolution of 4 cm-1. The background spectrum is collected from an unused area of the reflective support using the same aperture configuration and number of scans as used for the paint sample.
ATR | 7.8.5 | Section Title
ATR | 7.8.5.1 | Informational Text | ATR microscope objectives can be fitted with a silicon, ZnSe, diamond, KRS-5, or germanium internal reflecting crystal element. Different internal reflection elements (IREs) offer a variety of penetration depths and physical or chemical attributes. The sample is viewed using visible light and the area to be analyzed is centered in the field of view. The crystal is then placed in direct contact with the area of interest. Monitoring the single beam spectrum provides an indication of whether there is sufficient contact between the sample and crystal. Typically 64 to 512 scans are collected at a resolution of 4 cm-1. The number of scans collected for each sample can vary with size or type.
Sample Preparation Methods And Accessories | 8 | Section Title
Sample Preparation Methods And Accessories | 8.1 | Requirement | The method chosen for sample preparation depends on the size, nature, and condition of the specimen, as well as the particular FTIR technique/accessory that is employed for analysis. Use multiple preparation or analytical techniques, or both as necessary to analyze all layers and characteristics.
Transmission Techniques (Main Bench) | 8.2 | Section Title
Transmission Techniques (Main Bench) | 8.2.1 | Requirement | Samples prepared for analysis by main bench transmission techniques shall be thin enough to allow infrared radiation to pass through without over-absorption. For transmission data that are viewed in absorbance, the sample optimally should be thin enough to produce a maximum absorbance of 1 absorbance unit. For transmission data viewed in % transmittance, spectral peaks ideally should not fall below 10 % T. This typically requires thicknesses of approximately 5 to 10 μms.
Transmission Techniques (Main Bench) | 8.2.2 | Recommendation | The separation and analysis of individual layers is recommended in order to determine chemical composition and detailed spectral characteristics of each layer. This can be achieved by microtomy or by hand using a sharp blade while observing the evidence under a stereomicroscope.
Transmission Techniques (Main Bench) | 8.2.3 | Informational Text | Sample preparation techniques which can be employed for transmission analysis using the main bench include a thin peel stretched over an aperture, an alkali halide pellet (for example, potassium bromide), or a diamond cell.
Transmission Techniques (Main Bench) | 8.2.4 | Informational Text | Thin peels of each layer can be placed on a glass microscope slide and compressed with the flat beveled surface of a scalpel blade, a roller bearing tool, or other equivalent technique, and then placed directly over a small masking aperture for analysis. Given the small sample sizes encountered, a beam condenser is typically employed.
Transmission Techniques (Main Bench) | 8.2.5 | Recommendation | Potassium bromide (KBr) pellets are made by grinding a small aliquot of the individual paint layer with dry spectroscopic grade KBr using a mortar and pestle. The powder is then transferred to a die maker and a press is used to generate the pellet. As KBr is deliquescent, the pellets, KBr powder and die maker should be stored in an oven or desiccator. Given the small sample sizes encountered, one to three millimeter diameter pellets mounted in a beam condenser are typically employed.
Transmission Techniques (Main Bench) | 8.2.6 | Informational Text | The diamond anvil cell is a useful sampling technique when the greatest spectral range is required and for laboratories that do not have a microscope accessory for their infrared spectrometer. It is possible to obtain spectra down to approximately 200 cm-1 with cesium iodide (CsI) optics, or down to 400 cm-1 with KBr optics. Both high- and low- pressure diamond cells can be used in conjunction with a beam condenser. Pliable and powdered samples are amenable to low-pressure diamond cells, while harder paints can require the use of high-pressure diamond cells. Low pressure diamond cells have the advantages of costing less, are compatible with an infrared microscope accessory and do not obscure spectral information in the 2400 to 1800 cm-1 region because they contain thinner diamonds than high pressure cells.
Transmission Techniques (Main Bench) | 8.2.7 | Informational Text | Diamond cells permit relatively simple sample preparation. The cell consists of two diamond windows, a holder for each diamond and a means of compression that achieves an appropriate sample thickness. There are a variety of designs available. The paint is simply placed on one of the diamond faces, the second diamond is positioned on top, and sufficient pressure is applied to form a film. This is normally done under a stereomicroscope to ensure uniform coverage of the diamond face.
Transmission Techniques (Main Bench) | 8.2.8 | Informational Text | For non-elastic paints, one diamond can be removed prior to analysis. This leaves the thin compressed film adhering to one of the diamond faces and avoids diffraction fringes in the recorded spectrum resulting from the two parallel diamond faces.
Transmission Techniques (Main Bench) | 8.2.9 | Informational Text | The diamond cell technique is essentially non-destructive because of diamond’s inertness. After analysis, the paint can be recovered uncontaminated from the diamond face using a scalpel blade or other suitable tool.
Transmission Techniques (Main Bench) | 8.2.10 | Requirement | The chief drawback of this technique is that a larger sample size is required than with a microscope accessory. This can be particularly significant when examining small paint chips that have multiple layers because extraction of sufficient paint from interior layers can represent a significant challenge. Furthermore, each sample shall be placed in the cell and removed prior to mounting the next sample, unlike the microscope method where multiple samples can be placed on the support material and analyzed sequentially.
Internal Reflection Techniques (Main Bench) | 8.3 | Section Title
Internal Reflection Techniques (Main Bench) | 8.3.1 | Informational Text | Attenuated Total Reflectance (ATR) spectroscopy can be used to analyze exposed paint layers. Although ATR spectroscopy accessories are available in both multi-reflection (macro scale) and single reflection (micro scale), only single reflection accessories will be discussed because they are more suitable for forensic paint examination. In contrast to transmission methods, ATR methods require virtually no sample preparation when examining an already exposed surface. In some instances ATR methods can lend themselves to conducting the examination in situ. Since ATR is a surface technique it is necessary to remove any extraneous material from the area to be examined.
Internal Reflection Techniques (Main Bench) | 8.3.2 | Informational Text | The bench ATR (single reflection) accessory uses an internal reflection element and, or optics to condense the beam onto a spot-sized sampling area. The crystal is mounted horizontally in a purged box within the sample compartment. The sample is placed in contact with the center of the crystal, and an adjustable piston or is used to apply sufficient force to ensure contact. Mirrors focus the beam such that the infrared light reflected by the sample is directed to the detector.
Internal Reflection Techniques (Main Bench) | 8.3.3 | Informational Text | ATR techniques are subject to inter-sample variations resulting from differences in area of contact of the sample with the IRE surface. Being a reflection technique, ATR spectra are prone to the same type of spectral distortions as noted for reflection spectroscopy.
FTIR Microscope Accessory | 8.4 | Section Title
FTIR Microscope Accessory | 8.4.1 | Informational Text | The use of a microscope accessory is preferred for very small samples and has several advantages over bench techniques. Spectra can be obtained from flattened paints as small as 10 to 20 μm in diameter. The microscope attachment permits the sequential analysis of multiple samples placed on an appropriate support material. The method also affords the advantage of optical viewing and choosing specific regions of interest in either heterogeneous samples or those having varying thickness.
FTIR Microscope Accessory | 8.4.2 | Requirement | Although it is a popular IR accessory for paint analysis, the microscope attachment has some disadvantages. Unlike DTGS detectors, the MCT detectors used in microscope accessories require cooling with liquid nitrogen to minimize electronic noise and achieve greater sensitivity. There is a compromise to be considered between sensitivity and spectral range with these detectors. Narrow band detectors that cut off in the 700 cm-1 range are more sensitive than the broad band detectors that cutoff in the 450 cm-1 range. Care shall be taken with very small samples as the use of small measurement apertures can limit longer wavelengths (smaller wavenumbers) radiation from reaching the detector. Heterogeneity issues are also more pronounced when using very small apertures.
FTIR Microscope Accessory | 8.4.3 | Informational Text | The chief drawback of this technique is its limited spectral range. Microscope optics and detectors have a cutoff of approximately 450 cm-1 as contrasted to the diamond anvil cell/beam condenser method with extended range optics affording a lower cutoff of approximately 200 cm-1. The lower range can be advantageous in the classification and comparison of inorganic pigments.
Transmission Measurements (Microscope) | 8.4.4 | Section Title
Transmission Measurements (Microscope) | 8.4.4.1 | Requirement | Transmission measurements are commonly used as they generate spectra with fewer artifacts. However, transmission methods generally entail more extensive sample preparation than reflection techniques. The paint shall be thin enough to avoid over-absorption.
Transmission Measurements (Microscope) | 8.4.4.2 | Requirement | Analysis of individual layers is required in order to spectrally characterize each layer in a multi-layered paint. For transmission FTIR spectroscopy with a microscope accessory, individual layers of multi-layered paints can be analyzed either as thin peels of each layer or as individual layers in a thin cross section of the intact chip. In comparative analysis, variations of sample thickness can generate spectral artifacts.
Transmission Measurements (Microscope) | 8.4.4.3 | Informational Text | The paint layers can be separated by hand using a sharp blade while observing the sample under a stereomicroscope. Thin peels of each layer can be placed on a glass microscope slide and compressed with the flat beveled surface of a scalpel blade, a roller bearing tool, or other suitable technique. The sample obtained can then be placed either directly over a small masking aperture or on an appropriate salt plate for analysis; a thickness on the order of 5 μm is desired.
Transmission Measurements (Microscope) | 8.4.4.4 | Informational Text | When selecting an appropriate infrared support material (for example, salt plate), several factors should be taken into consideration including cost, availability, environmental sensitivity, transparency of the material to infrared radiation, and the durability of the material.
Transmission Measurements (Microscope) | 8.4.4.5 | Informational Text | The low pressure diamond cell can also be used as a support medium under the FTIR microscope. The paint is simply placed on one of the diamond faces, the second diamond is positioned on top, and sufficient pressure is applied to form a film. This is normally done under a stereomicroscope to ensure uniform coverage. For non-elastic paints, one diamond is typically removed prior to analysis. This leaves the thin compressed film adhering to one of the diamond faces and avoids diffraction fringes in the recorded spectrum resulting from the two parallel diamond faces. The sampling technique is essentially non-destructive because of diamond’s inertness. The paint sample can be recovered uncontaminated from the diamond face using a scalpel blade or other appropriate tool.
Transmission Measurements (Microscope) | 8.4.4.6 | Requirement | The microscope attachment also affords FTIR analysis of multi-layered paints without physical separation of the layers. A thin cross section with intact layer structure can be cut by hand using a sharp blade or with a microtome (thicknesses on the order of 5 μm are typical). The cross sectioned sample is either flattened on a microscope slide and placed on a support material or flattened directly on the support material. If the sample is used for analysis by another technique (for example, elemental analysis), contributions from the support material can be detected. Individual layers can then be analyzed after observing with visible light and centering the sample in the field of view, delineating areas of interest using the microscope’s aperture controls. Care shall be taken to avoid contributions from adjoining layers.
Transmission Measurements (Microscope) | 8.4.4.7 | Informational Text | Cross section analysis has the advantage of viewing the sample optically and then selecting the specific region for analysis, thus permitting rapid analyses of individual layers. It also permits analyses of specific regions of inhomogeneous materials and compositional mapping. However, it typically requires smaller target apertures which can result in diffraction effects, heterogeneity concerns, and signal to noise constraints.
Reflection Measurements (Microscope) | 8.4.5 | Section Title
Reflection Measurements (Microscope) | 8.4.5.1 | Informational Text | If samples are compressed directly on metal, or a glass slide made of infrared light reflecting architectural glass (low e-glass), the microscope’s reflection mode can be used to produce double-pass transmission spectra. The technique is sometimes referred to as “transflection” or “reflection/absorption.” Some wavelength maxima shifts can be observed in intense absorption bands. Transflection samples need to be approximately half the thickness of that which is optimal for transmission measurements.
Reflection Measurements (Microscope) | 8.4.5.2 | Informational Text | The FTIR microscope can also be used in the specular reflection mode, but in most cases, transmission methods are preferred. Refractive index dispersion, and differences in infrared absorption for different wavelengths, gives rise to distortions in specular reflection spectra. Specular reflection spectra are not absorbance spectra and cannot be compared to transmission spectra. Also, most of the reference data of coatings, binders, pigments and additives consist of transmission spectra. Furthermore, being surface analysis techniques, inconsistencies in the preparation of sample surfaces can present problems in detailed comparisons. Additionally, when analyzing individual layers in cross section and using the requisite small apertures, signal-to-noise constraints are even greater than those encountered in transmission analyses. Refer to 7.8.3.1 for further detail.
Internal Reflection Techniques (ATR Microscope Objective) | 8.4.6 | Informational Text
Internal Reflection Techniques (ATR Microscope Objective) | 8.4.6.1 | Informational Text | ATR objectives are available for infrared microscope accessories. The technique requires little or no sample preparation and is non-destructive.
Internal Reflection Techniques (ATR Microscope Objective) | 8.4.6.2 | Informational Text | An ATR microscope objective can be used to analyze exposed paint layers. In some instances ATR methods can lend themselves to conducting the examination in situ. Since ATR is a surface technique it can be desirable to analyze any existing surface material of interest and then remove the surface material from the area and reanalyze the area.
Internal Reflection Techniques (ATR Microscope Objective) | 8.4.6.3 | Recommendation | The analysis of thin surface smears can result in contributions from substrate material or underlying layers. If the substrate material does not transmit infrared radiation, such as metal or glass, the ATR spectra appear more like transmission spectra the thinner the sample becomes. Hence, ATR spectra of thin samples (on the order of 1 to 2 microns) should not be directly compared to ATR spectra of thicker ones.
Internal Reflection Techniques (ATR Microscope Objective) | 8.4.6.4 | Informational Text | ATR is subject to inter-sample variations resulting from variations in pressure on the sample surface and variations in surface contact areas. Also, intra-sample variations can result from sample heterogeneity. Being a reflection technique, it is somewhat prone to the same type of spectral distortions noted for reflection spectroscopy.
Performance Checks | 9 | Section Title
Performance Checks | 9.1 | Requirement | The instrument shall be checked to ensure it is operating properly. Results of the check shall be documented. Before any performance checks, the instrument shall be thermally stable. It is recommended that the system be left on, or in a stand-by mode, as continuous operation is better for performance, stability, and prolonging the lifetime of the IR-source. For detectors that require liquid nitrogen, the detector takes approximately twenty minutes to cool and stabilize.
Performance Checks | 9.2 | Recommendation | Performance checks should be conducted at least once a month (or before use if used less frequently) depending upon how often the instrument is used, or the laboratory’s protocol. It is recommended that the built-in instrument test from the manufacturer be used for the performance checks, if available. Instrument performance tests should include evaluation of conditions such as wavenumber accuracy, signal-to-noise ratio, signal strength, and 100 % line. Performance tests recommended by the manufacturer or those outlined in Practice E1421 can also be used. A performance check should also be used for FTIR accessories.
Classification, Comparison, and Interpretation | 10 | Section Title
Classification, Comparison, and Interpretation | 10.1 | Informational Text | Binder classification of commonly encountered coatings is based on the interpretation of characteristic infrared absorption bands. Similarly, some pigments and additives can be identified.
Classification, Comparison, and Interpretation | 10.2 | Informational Text | Classification of a coating can be achieved by evaluating the absorption bands present in the spectrum with respect to band position, band shape and band intensity. After evaluating the absorption bands, interpretation and classification of a spectrum can be accomplished through comparison of the collected data to spectra of reference materials, use of flow charts and published findings. Spectral libraries of known materials can also be used to characterize the binder, additives, or pigments, or a combination thereof, present in the paint.
Classification, Comparison, and Interpretation | 10.2.1 | Recommendation | There are many sources of flow charts and functional group frequency charts. Flow charts are easy to use and offer an examiner a place to start in classifying paints. However, classification should not be solely based on comparison to flow charts. The spectrum should be compared to representative reference spectra before a classification is made. Moreover, new formulations may not be represented in existing paint classification systems.
Classification, Comparison, and Interpretation | 10.3 | Informational Text | Binder classification can be hindered by the presence of certain pigments. If the paint binder is soluble, a micro extraction can be performed to isolate the binder for analysis. If the paint is not soluble, then an alternative analytical technique, such as pyrolysis-gas chromatography, can be used to elucidate the classification of the binder.
Classification, Comparison, and Interpretation | 10.4 | Requirement | Binder classification can be difficult in contaminated paints and smears. Contributions from substrate or other co-mingled materials shall be considered. Particle picking can be used to obtain spectra which are suitable for interpretation. Spectral subtraction can also be used to assess the presence of co-mingled materials.
Classification, Comparison, and Interpretation | 10.5 | Informational Text | Classification of the binder type, pigments, or additives, or a combination thereof, can assist in determining the provenance and end use of the coating. The classification of the binder can differ from the manufacturer’s designation due to changes in polymer chemistry or component migration that occur during curing. Also, trade names given to the coating by the manufacturer may not reflect the actual chemistry of the paint. This conflict in designation results from the use of marketing trade names and historical designations that are technically inaccurate or incomplete.
Classification, Comparison, and Interpretation | 10.6 | Informational Text | Comparison of known and questioned evidence can be conducted with both spectra displayed in transmittance or absorbance, or both, although certain information can be observed more readily in one format or the other.
Classification, Comparison, and Interpretation | 10.6.1 | Informational Text | There are a number of factors to consider when assessing whether or not spectra can be distinguished from one another: the presence or absence of absorption bands, their positions, shapes, and relative intensities.
Classification, Comparison, and Interpretation | 10.6.1.1 | Recommendation | Characteristic absorption bands present in one spectrum should be present in the comparison spectrum. The position of the absorption bands should agree with each other and is somewhat dependent on the shape of the absorption band. A rule of thumb is that the positions of corresponding peaks in two or more spectra be within ± 5 cm-1. For sharp absorption peaks one should use tighter constraints. One should critically scrutinize the spectra being compared if corresponding peaks vary by more than 5 cm-1. Replicate collected spectra can be used to determine reproducibility of absorption position.
Classification, Comparison, and Interpretation | 10.6.1.2 | Requirement | The absorption bands should have comparable relative intensities and shapes for the spectra being compared. If there is variation between spectra, acquire additional spectra if possible to determine if the variation is reproducible. (1) If there are notable differences among spectra of a single sample, the collection of additional spectra can be used to assess the range of variation. (2) If differences are noted between questioned and known items, the collection of additional spectra can be used to demonstrate whether the differences are repeatable and therefore meaningful.
Classification, Comparison, and Interpretation | 10.7 | Informational Text | One of three conclusions can be reached after evaluating and comparing the known and questioned spectra.
Classification, Comparison, and Interpretation | 10.7.1 | Informational Text | Spectra are dissimilar if they contain one or more meaningful differences.
Classification, Comparison, and Interpretation | 10.7.2 | Informational Text | Spectra are indistinguishable if they contain no meaningful differences.
Classification, Comparison, and Interpretation | 10.7.3 | Informational Text | A spectral comparison is inconclusive if sample size or condition precludes a decision as to whether differences are meaningful.
Spectral Libraries | 11 | Section Title
Spectral Libraries | 11.1 | Recommendation | Several infrared spectral libraries and databases relating to paint are available. The Federation of Societies for Coatings Technology offers a library in both printed and digital formats. It is a compilation of the infrared spectra of the various chemical compounds commonly found in coatings. Other libraries relating to polymers, pigments, and additives found in coatings are marketed by various companies. The ability to search the library of spectra and compare them to the spectrum of a specific paint layer can assist in the identification of the chemical components in that layer. The analyst should be certain that the library spectra are collected using the same method as that used in the analysis of the evidence.
Paint Data Query Database (PDQ) | 12 | Section Title
Paint Data Query Database (PDQ) | 12.1 | Informational Text | For forensic purposes, one of the most comprehensive compilations of OEM (original equipment manufacture) automotive paint information is the Paint Data Query (PDQ) database. PDQ is used to aid in the identification of make, model, and year of an unknown vehicle, to assess the relative significance of a paint system when conducting a paint comparison, and to stay current with automotive paint trends. PDQ consists of a text-based database and spectral libraries. It was created in the early 1970s and is maintained by the Royal Canadian Mounted Police Forensic Laboratory Services. This database is only available to law enforcement agencies through contractual agreement.
Paint Data Query Database (PDQ) | 12.1.1 | Informational Text | The PDQ text-based database consists of two parts; the source-based information and the paint layer information. The source-based information includes topics such as production plant, make, model, and year for each paint specimen in the database. The paint layer information includes the layer system (number, color, sequence, etc.) and the chemical composition of each layer.
Paint Data Query Database (PDQ) | 12.1.2 | Informational Text | The PDQ spectral libraries include the infrared spectra of individual paint layers.
Paint Data Query Database (PDQ) | 12.1.3 | Informational Text | To aid in the identification of make, model, and year of an unknown vehicle, the data acquired from an evidentiary paint system is coded according to PDQ format and is searched against the text-based portion of the database. The spectral library is also available for searches and comparisons to the data obtained from evidentiary paints to further narrow the range of possible sources.
Paint Data Query Database (PDQ) | 12.1.4 | Informational Text | The database can also serve to assist in assessing the relative significance of a correspondence between a questioned and known OEM paint system. For example, accessing the database provides information as to how many other makes and models of vehicles in the database used the same paint system.
Paint Data Query Database (PDQ) | 12.1.5 | Informational Text | The database is also a source of information to keep the forensic paint examiner aware of current trends in OEM automotive coating systems. These include changes in technology such as paint compositions and paint layer structures.
Paint Data Query Database (PDQ) | 12.2 | Requirement | PDQ is intended to be a source-based database, not a population-based database. In other words, PDQ does not contain information on finish systems of all vehicles in existence. This shall be taken into consideration whenever the database is used.
Documentation | 13 | Section Title
Documentation | 13.1 | Recommendation | Case notes should include a copy of all of the instrumental data that was used to reach a conclusion. All hard copies should include a unique item designation, the operator's name or initials, and the date of analysis.
Documentation | 13.2 | Recommendation | Case notes should also include a description of the evidence analyzed by IR, the method of sample preparation, the analytical instrumentation used, and its operating parameters.
Documentation | 13.3 | Informational Text | See SWGMAT's Trace Evidence Quality Assurance Guidelines for further requirements.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics